## Slide 1: Unique Resolution Standard Kit Fabricated by Proton Beam -Facilitates Accurate & Fast Beam Size Measurements
Prof Jeroen van Kan

## Slide 2: Background Inventor
Prof Jeroen van Kan
Director of Centre for Ion Beam Applications a collaborative Center to the IAEA
Vice President of the Microscopy Society Singapore
Deputy Director for the Engineering Science Programme, NUS
Member of the editorial board of Microelectronic Engineering
Department of Physics
National University of Singapore (NUS)
E-mail: Phyjavk@nus.edu.sg
2

## Slide 3
Resolution Standard Kit
[1] https://www.tedpella.com/calibration_html/SEM_Resolution_Test_Specimens_Tin_on_Carbon.aspx
[2] Y. Yao et al. Improved beam spot measurements in the 2nd generation proton beam writing system (2013)
Overview
Used for calibrating and characterising spatial resolution and scan dimensions of imaging systems
Requirements for charged-particles imaging include high spatial resolution, precise beam spot size measurement with low uncertainty
Two kinds of resolution standards:
  Spherical with sizes ranging from a few nm to hundreds of nm
  Grid with periodic patterns
Commercial spherical tin balls [1]
Nickel grids [2]
3

## Slide 4
Interactions of MeV Light Ions With Matter
4
Advantages of Our Standards
CIBA’s Proton Beam Writing (PBW) able to produce grids with straight side walls [5]
Side wall angle ~89.9°, much better than industry standard of ~88.5° [1]
2 MeV
Protons
50 keV
Heavy Ions
50 keV
Electrons
60 µm
Different particles interacting with medium [5]
Resolution Determination Methods
Edge Scanning scans edges of features, resulting line profile fitted using mathematical function to determine resolution [3]
  Spherical RS relies on curvature at the side; Grids with straight side-walls may perform better, eliminating error introduced by curvature of sphere
Gap Scanning scans gaps between adjacent features [4]
  No fixed gap sizes for spherical RS, best resolution depends on smallest gap available; Edge Scanning with Grids only utilises side-wall straightness, eliminating physical limits of gaps
Sidewall angle
[3] J. Swindells et al. Modelling PSF of scanning electron microscopes for image restoration (1998)
[4] V. András et al. On the Sub-Nanometer Resolution of Scanning Electron and Helium Ion Microscopes (2009)
[5] F. Watt et al. Proton beam writing (2007)

## Slide 5
Depth (nm)
Udalagama et al. PRB (2009)
Reduced Proximity Effects  High Density & Fine Structures (Analysis & Lithography)
Depth can be tuned  High aspect ratio, Multi Level structures & 3D Analysis
Straight trajectory  Smooth Side Walls & Well resolved features
High sensitivity – 100 fold higher than electrons (Single Ion Imaging)
Spatial Energy
Deposition Profile
using DEEP code
P-beam writing is potentially superior compared to E-beam!
100 keV Electrons
500 keV Protons
E-Beam ↔ P-Beam
Energy Deposition in 1 µm PMMA (Delta Rays)
5
Electron Beam & Proton Beam Induced Delta Rays or Secondary Electrons

## Slide 6
Three-dimensional nanolithography using proton beam writing
J.A. van Kan, et al Appl. Phys. Lett., 83 (2003) 1629
50 nm lines in PMMA
60 nm
10 mm
Thick SU-8
Proton Beam Writing of Three-Dimensional Nanostructures in HSQ
J.A. van Kan, et al
Nano Letters 6 (2006) 579
PMMA Requires 3.5 eV bond breaking
HSQ requires 4 & 8 eV cross linking
These electron are used for imaging
19 nm
100 nm thick HSQ
Microsyst Technol, 2014
2014
PBW: High Aspect Ratios in Resist
6

## Slide 7
Proton Imaging: IBAC & STIM
Ni Grid 1 µm thick
IBAC (Ion Beam Absorbed Current)
STIM (Scanning Transmission Ion Microscopy)
35 x 101 nm @ (1.5 MeV H2+ 1 pA)
IBAC
STIM
X - 35 nm
Y - 101 nm
7

## Slide 8
Proton Imaging: STIM & Electron Detection
Ni Grid 1 µm thick
On/Off Axis STIM (Scanning Transmission Ion Microscopy)
Proton Induced Electron Imaging
15 x 36 nm @ (2 MeV H2+ ~ fA)
On Axis STIM
Off-axis STIM
Electron Detection
Beam Size: X - 15 nm
Y - 36 nm
8

## Slide 9
STIM imaging of Ni grid
Detector
IBIC (Ion Beam Induced Current) protons travel through 1 um on Ni before reaching PN junction
Sub 10 nm resolution
E
Counts
Si PN junction
1 µm
Ni grid
Ni Grid
1 µm thick
9
Ultra Microscopy 2023
Proton Microscopy at Sub-10 nm

## Slide 10
Result Was Proven
10
Our technology was published by many established scientific journals [1-3]
Proven Record
  We have been using the same grid for more than 10 years without degrading performance in CIBA PBW system for setting up all kinds of Proton Microscopy experiments for us & customers
Commercial Usage-Customers Validation
  Some customers have bought our kits
  Gave positive feedback that our kits have more superior function than others . Which enables them to have better performance
[1] J. A. van Kan et al. Fabrication of a free standing resolution standard for focusing MeV ion beams to sub 30 nm dimensions (2005)
[2] Y. Yao et al. Improved beam spot measurements in the 2nd generation proton beam writing system (2013)
[3] S. Qureshi et al. Automated alignment and focusing system for nuclear microprobes (2019)
[4] Y. X. Dou et al. Breaking the 10nm barrier using molecular ions in nuclear microprobes (2023)

## Slide 11
Conclusions:
Resolution Standard Kit via Proton Beam
  Potential
  Consistent production of Resolution Standard Kit with proven high resolution capability
  Tried-and-tested SOP for production
  Proven track record of consistent small beam spot sizes with heavy usage over a decade
  Commercially available
11

## Slide 12
Support from:
NRF, MOE, NUS (Singapore),
US air force
Thank You
12

## Slide 16
Historic beam spot size improvement
Normally, a nuclear microprobe has a different resolution in two lateral directions (i.e. rectangular profile).
Here, the definition of resolution is the square root of the product of resolutions in X & Y axes.
CIBA has been at the forefront of resolution improvements with MeV proton beams for most of the 21st century!
*Reported 1-D resolution
Dou, Y. et al. Breaking the 10 nm barrier using molecular ions in nuclear microprobes Ultramicroscopy 253, 113812 (2023)
Proton Microscopy at Sub-10 nm
16

## Slide 17
CIBA Ion Beam Facility
17

## Slide 18
Multi disciplinary research utilising high energy light ion beams
Biophysics
IC Failure
Analysis
Materials
characterisation
Microphotonics
Nano
engineering
Micro/Nano
fluidics
Silicon
micromachining
Nano-replication
Tissue-engineering
Proton Beam Writing Line
18

## Slide 19
1000 nm
100 keV Electrons
500 keV Protons
Radial Distance (nm)
2 nm
1 nm
Normalized Deposited Energy
1st Gen-e
2nd Gen-e
3rd Gen-e
4th Gen-e
All (e + P)
Primary P
PRB (2009)
Proton Interaction With Material at Nano Scale
19
Notes: 19

## Slide 20
Neg or Positive Resist
Ni
Seed layer
Chemical
development
Proton beam
exposure
Electroplating
Res Kit
Strip
SU-8 or PMMA
Ni structure
Crosslink
Negative resist
Chain scission
Positive resist
Insoluble for
developer
Remove
exposed resist
PMMA mold for plating
Ni
H: 5 mm
Au
35 nm
700 nm
Proton Beam Writing
20
SEM image of grid showing orthogonality (left) and smoothness (right) of grid bars [3]
Notes: 20